# Circuit Designs ESP32S3 Part II

# Question: When the ALERT/RDY pin is signaled, how do we know which ADC channel triggered it? OR do we assume that all 4 channels are synchronized?

When the ALERT/RDY pin signals, it simply tells you that the most recently requested conversion has finished. It does not provide any channel identification bits or metadata, nor are the channels synchronized.

Because the ADS1115 contains only one internal ADC core, it handles channels sequentially. You must use a deterministic state machine in your code to track which channel triggered the interrupt.

## How the State Machine Tracks Channels

Because the ADS1115 operates like a conveyor belt, your code must always stay one step ahead. The sequence follows a strict Pipeline Pattern:

[Time Line]  -------------------------------------------------------------------->

Hardware:   |  Converting Ch 0  |  Converting Ch 1  |  Converting Ch 2  | ...

Interrupt:                      | (ALERT Fires!)    | (ALERT Fires!)    |

Firmware:                       | Read Ch 0 Result  | Read Ch 1 Result  |

| Config MUX to 1   | Config MUX to 2   |

- Step 1: You tell the ADS1115 to sample Channel 0.

- Step 2: The ALERT pin drops low. You know this data belongs to Channel 0 because your code explicitly asked for it in Step 1.

- Step 3: Your ISR wakes up your task. You read the data register (saving it as Channel 0) and immediately write to the ADS1115 configuration register to start the conversion for Channel 1.

- Step 4: The state machine increments its internal index tracker (channel_index = 1).

When the ALERT pin drops low again, your state machine checks its tracker, sees 1, and knows the incoming data belongs to Channel 1.

## The "First Sample" Pitfall (Crucial for 860 SPS)

When changing channels, the single internal ADC core requires a settling time. The very first ALERT signal after switching the multiplexer config register does not contain the new channel's data—it contains the final conversion of the previous channel that was already in flight.

To achieve an accurate 860 SPS sweep across 4 channels, you cannot use "Continuous Mode" seamlessly across multiple channels. Instead, you must switch the chip into Single-Shot Mode.

The corrected logic for your high-speed loop operates like this:

| Step | State Machine Track | Hardware Action | Firmware Action on Interrupt |
| --- | --- | --- | --- |
| 1 | Track = 0 | Chip samples Channel 0 | ALERT fires → Read data into Channel 0 structure. |
| 2 | Switch Track to 1 | Command Single-Shot on Ch 1 | Chip begins conversion. Firmware sleeps. |
| 3 | Track = 1 | Chip samples Channel 1 | ALERT fires → Read data into Channel 1 structure. |
| 4 | Switch Track to 2 | Command Single-Shot on Ch 2 | Chip begins conversion. Firmware sleeps. |

## Code Fix for the Acquisition Task

Modify your adc_acquisition_task in main.c to use Single-Shot mode (0x8000 bit flipped) so the ALERT pin matches the channel you expect:

c

// Mask changes: Bit 8 is set to 1 for Single-Shot conversion mode static const uint16_t ads_mux_config[4] = { 0xC4E8, // Channel 0, Single-Shot, 860 SPS, Alert Active Low 0xD4E8, // Channel 1, Single-Shot, 860 SPS, Alert Active Low 0xE4E8, // Channel 2, Single-Shot, 860 SPS, Alert Active Low 0xF4E8 // Channel 3, Single-Shot, 860 SPS, Alert Active Low }; void adc_acquisition_task(void *pvParameters) { uint8_t tracking_channel = 0; adc_sample_t current_batch = {0}; uint16_t raw_reading = 0; // Trigger the very first single-shot conversion on Channel 0 ads1115_write_reg(ADS1115_CFG_REG, ads_mux_config[tracking_channel]); while (1) { // Wait for the ALERT pin interrupt semaphore if (xSemaphoreTake(ads_rdy_semaphore, portMAX_DELAY) == pdTRUE) { // Read conversion register (Guaranteed to be tracking_channel data // in Single-Shot mode) if (ads1115_read_reg(ADS1115_CONV_REG, &raw_reading) == ESP_OK) { current_batch.channels[tracking_channel] = (int16_t)raw_reading; } // Move state machine tracking to the next channel index uint8_t next_channel = (tracking_channel + 1) % 4; if (tracking_channel == 3) { current_batch.timestamp = xTaskGetTickCount() * portTICK_PERIOD_MS; xQueueSend(adc_queue, &current_batch, 0); } tracking_channel = next_channel; // Trigger the next single-shot conversion immediately ads1115_write_reg(ADS1115_CFG_REG, ads_mux_config[tracking_channel]); } } }

Use code with caution.

I can help you round out this firmware. Would you like me to add the DS3231 RTC I2C polling code into this exact loop so your SD card logs use a real-world clock timestamp instead of an internal processor uptime millisecond counter?

# Yes, please do add the DS3231 RTC I2C polling code into this exact loop. This is the required action that must be taken for a useful ECG trace being sampled.

To capture a useful ECG trace at 860 SPS, your time-tracking architecture must be highly specialized.

Reading the DS3231 RTC over I2C takes roughly 1 to 2 milliseconds. If you attempt to poll the RTC over I2C inside your high-speed ADC acquisition loop, you will stall the state machine and drop half of your ECG samples.

## The Hybrid Timestamping Solution

To fix this, the firmware uses a Hybrid Timestamping Engine:

- The Real-Time anchor: The slow DS3231 RTC is polled once per second on Core 1 (the SD Card thread) to establish an absolute wall-clock reference (e.g., 14:32:01).

- The Microsecond precision: The high-speed ADC thread on Core 0 uses the ESP32-S3's internal high-resolution hardware timer (esp_timer_get_time()). This records the exact microsecond each 4-channel ECG sweep occurs relative to the baseline.

- The Output: The SD card writer merges these two values. Your CSV output will contain an absolute RTC start time followed by microsecond-accurate steps, allowing you to reconstruct the ECG waveform with zero jitter.

## Updated Code Implementation (main.c)

c

#include <stdio.h> #include "freertos/FreeRTOS.h" #include "freertos/task.h" #include "freertos/queue.h" #include "freertos/semphr.h" #include "driver/i2c.h" #include "driver/gpio.h" #include "esp_timer.h" #include "esp_vfs_fat.h" #include "sdmmc_cmd.h" #define I2C_MASTER_SDA_IO 4 #define I2C_MASTER_SCL_IO 5 #define I2C_MASTER_NUM I2C_NUM_0 #define I2C_MASTER_FREQ_HZ 400000 #define ADS1115_I2C_ADDR 0x48 #define ADS1115_CONV_REG 0x00 #define ADS1115_CFG_REG 0x01 #define DS3231_I2C_ADDR 0x68 #define DS3231_TIME_REG 0x00 #define ADS_ALERT_GPIO 6 #define ESP_INTR_FLAG_DEFAULT 0 typedef struct { uint64_t micro_timestamp; // Microsecond accuracy for ECG spacing int16_t channels[4]; } adc_sample_t; typedef struct { uint8_t second; uint8_t minute; uint8_t hour; } rtc_time_t; static QueueHandle_t adc_queue = NULL; static SemaphoreHandle_t ads_rdy_semaphore = NULL; // Global wall clock shared safely via thread updates static volatile rtc_time_t global_wall_clock = {0}; static const uint16_t ads_mux_config[4] = { 0xC4E8, // Ch 0, Single-Shot, 860 SPS, Alert Active Low 0xD4E8, // Ch 1, Single-Shot, 860 SPS, Alert Active Low 0xE4E8, // Ch 2, Single-Shot, 860 SPS, Alert Active Low 0xF4E8 // Ch 3, Single-Shot, 860 SPS, Alert Active Low }; // Helper: Convert BCD format from DS3231 to regular integers static uint8_t bcd_to_dec(uint8_t val) { return ((val / 16) * 10) + (val % 16); } static esp_err_t ds3231_read_time(rtc_time_t *time) { uint8_t reg = DS3231_TIME_REG; uint8_t data[3]; esp_err_t err = i2c_master_write_read_device(I2C_MASTER_NUM, DS3231_I2C_ADDR, &reg, 1, data, 3, pdMS_TO_TICKS(10)); if (err == ESP_OK) { time->second = bcd_to_dec(data[0] & 0x7F); time->minute = bcd_to_dec(data[1] & 0x7F); time->hour = bcd_to_dec(data[2] & 0x3F); // 24-hour mode assumption } return err; } static esp_err_t ads1115_write_reg(uint8_t reg, uint16_t val) { uint8_t out_buf[3] = { reg, (uint8_t)(val >> 8), (uint8_t)(val & 0xFF) }; return i2c_master_write_to_device(I2C_MASTER_NUM, ADS1115_I2C_ADDR, out_buf, 3, pdMS_TO_TICKS(10)); } static esp_err_t ads1115_read_reg(uint8_t reg, uint16_t *val) { uint8_t in_buf[2]; esp_err_t err = i2c_master_write_read_device(I2C_MASTER_NUM, ADS1115_I2C_ADDR, &reg, 1, in_buf, 2, pdMS_TO_TICKS(10)); if (err == ESP_OK) { *val = (in_buf[0] << 8) ¦ in_buf[1]; } return err; } static void IRAM_ATTR ads_gpio_isr_handler(void* arg) { BaseType_t xHigherPriorityTaskWoken = pdFALSE; xSemaphoreGiveFromISR(ads_rdy_semaphore, &xHigherPriorityTaskWoken); if (xHigherPriorityTaskWoken) { portYIELD_FROM_ISR(); } } // CORE 0: High-Priority ECG Acquisition void adc_acquisition_task(void *pvParameters) { uint8_t tracking_channel = 0; adc_sample_t current_batch = {0}; uint16_t raw_reading = 0; ads1115_write_reg(ADS1115_CFG_REG, ads_mux_config[tracking_channel]); while (1) { if (xSemaphoreTake(ads_rdy_semaphore, portMAX_DELAY) == pdTRUE) { if (ads1115_read_reg(ADS1115_CONV_REG, &raw_reading) == ESP_OK) { current_batch.channels[tracking_channel] = (int16_t)raw_reading; } uint8_t next_channel = (tracking_channel + 1) % 4; if (tracking_channel == 3) { // High-precision microsecond hardware timer stamp for ECG baseline alignment current_batch.micro_timestamp = esp_timer_get_time(); xQueueSend(adc_queue, &current_batch, 0); } tracking_channel = next_channel; ads1115_write_reg(ADS1115_CFG_REG, ads_mux_config[tracking_channel]); } } } // CORE 1: Low-Priority RTC Syncing & SD Card Storage void sd_writer_task(void *pvParameters) { adc_sample_t log_data; TickType_t last_rtc_check = xTaskGetTickCount(); rtc_time_t local_clock_copy; // Seed the initial wall clock ds3231_read_time(&local_clock_copy); global_wall_clock = local_clock_copy; while (1) { // Safe RTC Polling: Only check the physical chip once every 1000ms if ((xTaskGetTickCount() - last_rtc_check) > pdMS_TO_TICKS(1000)) { if (ds3231_read_time(&local_clock_copy) == ESP_OK) { global_wall_clock = local_clock_copy; } last_rtc_check = xTaskGetTickCount(); } // Pull incoming high-speed ECG data from queue if (xQueueReceive(adc_queue, &log_data, portMAX_DELAY) == pdTRUE) { FILE* f = fopen("/sdcard/ecg_trace.csv", "a"); if (f != NULL) { // Formats as: HH:MM:SS,Microseconds,Ch0,Ch1,Ch2,Ch3 fprintf(f, "%02d:%02d:%02d,%llu,%d,%d,%d,%d\n", global_wall_clock.hour, global_wall_clock.minute, global_wall_clock.second, log_data.micro_timestamp, log_data.channels[0], log_data.channels[1], log_data.channels[2], log_data.channels[3]); fclose(f); } } } } void app_main(void) { // I2C, GPIO, and Semaphore Initialization steps remain identical to previous configuration i2c_config_t conf = { .mode = I2C_MODE_MASTER, .sda_io_num = I2C_MASTER_SDA_IO, .sda_pullup_en = GPIO_PULLUP_ENABLE, .scl_io_num = I2C_MASTER_SCL_IO, .scl_pullup_en = GPIO_PULLUP_ENABLE, .master.clk_speed = I2C_MASTER_FREQ_HZ, }; i2c_param_config(I2C_MASTER_NUM, &conf); i2c_driver_install(I2C_MASTER_NUM, conf.mode, 0, 0, 0); gpio_config_t io_conf = { .intr_type = GPIO_INTR_NEGEDGE, .pin_bit_mask = (1ULL << ADS_ALERT_GPIO), .mode = GPIO_MODE_INPUT, .pull_up_en = GPIO_PULLUP_ENABLE }; gpio_config(&io_conf); ads_rdy_semaphore = xSemaphoreCreateBinary(); adc_queue = xQueueCreate(100, sizeof(adc_sample_t)); // Expanded queue size for dense trace data gpio_install_isr_service(ESP_INTR_FLAG_DEFAULT); gpio_isr_handler_add(ADS_ALERT_GPIO, ads_gpio_isr_handler, NULL); // Pin high priority capture to Core 0, file management to Core 1 xTaskCreatePinnedToCore(adc_acquisition_task, "ECG_Acq", 4096, NULL, 10, NULL, 0); xTaskCreatePinnedToCore(sd_writer_task, "SD_Write", 4096, NULL, 4, NULL, 1); }

If you are interested, we can customize this further:

- Would you like to implement an analog signal filtering network (like a hardware RC low-pass filter) to clean up high-frequency noise before it reaches the ADS1115 inputs?

- Do you want to configure the ADS1115 Programmable Gain Amplifier (PGA) settings specifically to match the voltage range of your ECG front-end hardware?

# We will be using the AD8232 to acquire the signals from the electrodes, and the AD8232 does incorporate sufficient filtering for ECG monitoring, so that should be more than adequate (although at a later time, I might inquire about an objective analysis of the AD8232). Yes, please do show what needs to be added to handle the ADS1115 PGA.

To interface the AD8232 ECG front-end with the ADS1115 effectively, you must configure the ADS1115's internal Programmable Gain Amplifier (PGA) to match the specific voltage output profile of the AD8232.

## AD8232 Voltage Output Matching

The AD8232 scales and centers the microvolt-level electrical signals from heart tissue using an internal reference buffer. It operates under specific electrical bounds:

- Output Voltage Range: The AD8232's output signal swings between 0V and 3.3V (relative to system ground).

- The Baseline Center: When the patient is completely still, the baseline ECG trace rests right in the middle at approximately 1.65V.

Because the signal peaks at 3.3V, setting the ADS1115 PGA to its maximum sensitivity (such as ±0.256V or ±0.512V) will cause the signal to hard-clip, resulting in a flat-line trace. Conversely, setting it too wide (such as ±6.144V) reduces your dynamic resolution.

The optimal setting is PGA = 1, which establishes a Full-Scale Range (FSR) of ±4.096V. [1]

ADS1115 Full-Scale Range (±4.096V)

+4.096V --------------------------------------------------

|   AD8232 Max Peak (~3.3V)

|     /\      /\

|----/--\----/--\---- AD8232 Baseline Reference (~1.65V)

|   /    \  /    \

|__/______\/______\__ AD8232 Min Trough (~0V)

0.000V --------------------------------------------------

-4.096V --------------------------------------------------

While the AD8232 will never generate negative voltages, this setting ensures its entire 0–3.3V range fits safely within the ADC's sampling window without clipping, while utilizing over 80% of the available digital resolution.

## Modifying the Binary Register Configurations

To change the PGA setting to ±4.096V, you must modify Bits [11:9] inside the ADS1115 Configuration Register to 001.

Updating the configuration hex array in your code from X4EX to X2EX applies this gain adjustment across all four channels while maintaining your 860 SPS single-shot hardware alert settings:

c

// Updated Configuration Matrix for PGA = +/-4.096V (Bits [11:9] = 001)

static const uint16_t ads_mux_config[4] = {

0xC2E8, // Channel 0, Single-Shot, 860 SPS, FSR = +/-4.096V, ALERT Line Active

0xD2E8, // Channel 1, Single-Shot, 860 SPS, FSR = +/-4.096V, ALERT Line Active

0xE2E8, // Channel 2, Single-Shot, 860 SPS, FSR = +/-4.096V, ALERT Line Active

0xF2E8  // Channel 3, Single-Shot, 860 SPS, FSR = +/-4.096V, ALERT Line Active

};

Use code with caution.

## Converting Raw ADC Readings into Real Voltages

With the PGA set to ±4.096V, the ADS1115 maps its 16-bit signed internal register across a range of 65,536 discrete steps. Each step (or Least Significant Bit) equates to exactly 125 μV: [1, 2]

\(\text{LSB\ Size}=\frac{4.096\text{\ V}}{32,768\text{\ steps}}=0.000125\text{\ V}=125\,\mu \text{V}\)

To save actual millivolt values directly onto your SD card rather than uncalibrated raw binary steps, update the fprintf statement inside your sd_writer_task loop:

c

// Convert the raw int16_t reading into millivolts using the 125uV multiplier

float ch0_mv = log_data.channels[0] * 0.125f;

float ch1_mv = log_data.channels[1] * 0.125f;

float ch2_mv = log_data.channels[2] * 0.125f;

float ch3_mv = log_data.channels[3] * 0.125f;

// Save clean decimal millivolt metrics directly into your CSV log

fprintf(f, "%02d:%02d:%02d,%llu,%.3f,%.3f,%.3f,%.3f\n",

global_wall_clock.hour,

global_wall_clock.minute,

global_wall_clock.second,

log_data.micro_timestamp,

ch0_mv, ch1_mv, ch2_mv, ch3_mv);

When you are ready to expand on this setup, let me know if you would like to:

- Handle the AD8232's digital hardware lines, such as mapping the Leads-Off Detect (LO+ / LO-) pins to the ESP32-S3 to detect if an electrode detaches from the patient.

- Proceed with the objective analysis of the AD8232's analog filtering performance and noise performance at 860 SPS. [1]

# I have been considering how to handle Leads Off Detection on this ESP32S3 mini board platform in the light of scarce GPIO pins. I did come up with a scheme using two 3.3v quad OR gates that will trigger a single GPIO input when any of the Leads Off Detect lines go high. That way, it keeps the cost done. But, are there any other intermediary step that can generate a signal to actual specify which LO+ or LO- line went off on such a GPIO limited system?

Your idea of using logic gates to combine the flags is an excellent way to save precious pins! However, a basic quad OR gate gives you an "all or nothing" interrupt—you know a lead fell off, but you don't know which one.

Since you are already using a high-speed, 4-channel ADS1115 ADC and a shared I2C bus, you can actually identify exactly which lead disconnected without using any extra ESP32-S3 GPIO pins.

Here are the two best intermediary methods to achieve this:

## Option 1: The "Analog Resistor Ladder" Method (Zero Extra Pins)

Instead of feeding your quad OR gate output into a digital GPIO pin, you can feed a custom resistor mixer network directly into one of your unused ADS1115 channels (Channels 1, 2, or 3, assuming your main ECG signal is on Channel 0).

You can connect the LO+ and LO- lines to a single analog input through different resistor values to create a binary voltage divider:

3.3V Rail --------- [ 10kΩ ] ---------+--------- To ADS1115 Channel 1

|

LO+ Out ---- [ 10kΩ ] ----|<|---------+  (Diode isolates lines)

|

LO- Out ---- [ 20kΩ ] ----|<|---------+

Because the resistors have different values, each combination creates a unique voltage level when a lead goes high:

| Lead Status | Voltage at ADS1115 Pin | How Your ESP32-S3 Interprets It |
| --- | --- | --- |
| All Leads Connected | ~0.00 V | Everything is normal. |
| Only LO+ Disconnected | ~1.65 V | LO+ is disconnected. |
| Only LO- Disconnected | ~2.20 V | LO- is disconnected. |
| Both Disconnected | ~2.64 V | Both leads are disconnected. |

Why this fits your system: You are already sampling all 4 channels sequentially at 860 SPS. Channel 1 will automatically log the exact lead-off status in your CSV file on every single sweep without needing a single digital GPIO pin on the ESP32-S3.

## Option 2: The I2C I/O Expander Method ($0.50 Component)

If you want a pure digital solution without using up an ADC channel, replace your OR-gate chip with a small, cheap I2C I/O Expander like the PCF8574 or TCA9534.

- How to wire it: Connect the expander to your existing I2C lines (SDA / SCL). Connect your LO+ and LO- lines directly to its input pins. [1, 2]

- The interrupt: These expanders have an INT (Interrupt) pin. You can route this single pin into your existing OR-gate alert line, or route it directly to the ESP32-S3.

- How it works: When a lead falls off, the expander triggers your single interrupt pin. The ESP32-S3 simply reads a single byte over the I2C bus to see exactly which pin changed state.

## System Comparison

| Feature | Resistor Ladder (Option 1) | I2C Expander (Option 2) | Your Quad OR Gate |
| --- | --- | --- | --- |
| Extra ESP32 Pins | 0 | 0 (Shares I2C Bus) | 1 Digital Pin |
| Identifies Line? | Yes (Via Voltage Level) | Yes (Via Digital Bit) | No (Global Alarm Only) |
| PCB Space | Tiny (3 Resistors, 2 Diodes) | Small (1 SOIC-8/16 Chip) | Medium (1 SOIC-14 Chip) |
| Firmware Impact | None (Already sampling Ch 1) | Mild (Requires I2C read) | None |

If you want to try the Resistor Ladder approach, I can calculate the exact voltage thresholds and add the code logic to decode the values inside your existing sd_writer_task. Let me know how you want to design this!